Heartworm Testing

1.    Place 4 drops of blue conjugate solution into clear sample tube.

2.    Open SNAP Heartworm Test Kit.

3.    Using pipette provided, dispense 3 drops of blood sample into the blue solution in the clear sample tube.

4.    Cap the sample tube and recap the purple top tube and set aside.

 (Do not discard unused blood sample until test is fully run and a result is given.)

5.    Mix sample thoroughly by inverting it 3-5 times.

6.    Place testing device on a horizontal surface. Add the entire contents of the sample tube to the sample well, being careful not to splash the contents outside of the sample well. The sample will flow across the result window reaching the end of the result window in 30-60 seconds. Some sample may remain in the sample well.

7.    When color reaches the end of the result window, push activator firmly until it is flush with the device body.

8.    Start timer. Read test results at the end of the 8 minutes.

9.    Refer to test results diagram located on the clinic wall to interpret results.

10. Record results and testing date on animal's ticket.